## Document Summary
File | Year | Technology | Calibration Transfer Method | Type | Application | Type of Transfer | Notes
'Alamar Calibration_Transfer_Between_NIR_Diode_Array_and_F 2005.pdf' | 2005 | NIRS | PDS | Peer Reviewed academic paper | Apple Soluble Solids | Inter-instrument | model derived from FT instruments applied to diode array instruments. Found that standardizing poorer instruiments (diode array) to high performance was more effective than between poor instruments
Brouckaert calibration tramnsfer Raman lab to inline industrial 2017.pdf' | 2017 | Raman | S/B | Peer Reviewed academic paper | Detergent Composition | Inter-instrument | Transfering at-libne calibration model to use for in-line prediction. Used EMSC as a preprocessing step to standardise the spectra. Worked effectively, only test to fail was difference in reproduicibility between at line and inline#
'Bu pharmamanufacturing_com NIR Calibration Transfer 2008.pdf' | 2008 | NIRS | DS, PDS, S-W | Trade publication | General | Inter-instrument | focus on PAT, need to standardise output. Recommends nforward transfer (make model like instrument to be deployed on) rather than backward transfer (make new iinstruments look like original system)
'Dardenne Calibration Transfer in NIRS NIR news 2002.pdf' | 2002 | NIRS | S/B, Robust Calibration, S-W, PDS, Wavelet | Conference Proceedings | General | Inter-instrument | Discusses reasons for calibration transfer then different approaches. Uses Shenk-westerhaus to repesent spectral correction and guides reader to Tom Fearn's paper for other forms such as PDS and DS
'Dorantes Calibration set optimization and library transfer for soil carbon estimation 2022.pdf' | 2022 | NIRS | Sampling strategy | Peer Reviewed academic paper | Soil Organic Content | Sample Populations | Focus on sample strategies to optimise transferability, importance of exckuding outliers in calibration (so model only influenced by variations that have sufficicent statistical power). Representativeness is flagged as a critical factor for the calibration (coverage of the variatioon that will be encountered in real use). Local calibrations discussed, building models using samples expected to be similar in composition to the sample being measured. Wavelength selection discussed (only bretain regions that vary in relevant ways) (Note - this paper discusses RPD - ignore this, it is just an inverse R2). When tranfsering models between populations spiking the calibrations set with some samples from the application population improved transferability - this incorporates the new variation without needing a full set of samples for a new calibration.
'Eliarts calibration transfer benchtop and portable MIRS 2019.pdf' | 2019 | MIRS | S/B, PDS, Spiked Model | Peer Reviewed academic paper pre-proof | Cocaine classification | Inter-instrument | Benchtop to portable. Spiked ('mixed') model was best performance
'Fan long term evaluation of solids in apple NIR calibration transfer 2019.pdf' | 2019 | NIRS | S/B | Peer Reviewed academic paper | Apple Soluble Solids | Seasonal Variation | Simple bias/slope correction (S/B) improved RMSEP for subsequent year predictions from ~1% to 0.5% (latter being the typical within year RMSEP)
Fernandez_Ahumada NIR Calibrations from lab to process 2008.pdf' | 2008 | NIRS | S-W, PDS, spiked model | Peer Reviewed academic paper | Animal Feed | Inter-instrument | Transfer from pre-dispersive (grating monochromator) to post-dispersive (diode array) instruments. Main sample set used to calibrate on pre-dispersive instrument, a second data set was independent validation set for the calibration. A third set was measured on both systems to develop the mathematical corrections for aligning output. The validation set was also ran on the post-dispersive instrument to assess performance. Found that PDS and S-W quite comparable but found big improvement in transferability when they spiked the calibration set with spectra from instrument 2
'Feudale Transfer_of_multivariate_calibration_models 2002.pdf' | 2002 | NIRS (plus others mentioned) | Broad range | Peer Reviewed academic paper | General review article | Inter-instrument | Gave 3 main reasons for calibrations becoming invalid: 1. compositional inconsistency (bacth effects, seasonality etc)2. instrument response function 3. environmental changes (temperature, humidity etc). First step is to minimise differences e.g. eliminate instrumental response functions arising from non-sample sources (3.1 instrument matching). A second step is to calibration for as wide a range of expected variation as possible (3.2 global models). third is too update the old calibration with data covering the new variation (3.3 model updating). fourth is variable selection , selecting only variables that are important in sample and low variation otherwise (3.4 sensor selection). The next step is to standardise the calibration and the new data in some way. Options are 1. standardise the model cofficients (see Wang 1991) 2. standardise the spectral responses (DS, PDS, S-W) 3. standardise the predictions (S/B) 4. mapping data in a transformed space (FT, wavelet decomposition). Allso listed miscellaneous methods including non linear mapping using ANNs, maximum likelihood PCA, positive matrix factorization (both iterative methods to estimate what the spectra should be), Kalman and Wiener filters. Examples of these too few for author to evaluate. Suggest that Kalman filter mn+might be good for model updating in the face of drift. The next part of the process is selecting samples to develop and test the transfer process. It is essential to cover the full range of variation and preference should be given to smples that have a high leverage on the calibration. For some reason the authors leave signal processing to now, despite it being important in eliminating unwanted variability and therefore important in aligning instrument output. They mention smoothing (to minimise pixel to pixel noise variations such as shot noise, fixed pattern response etc), intensity normalisation, offset normalisation (via MSC and finite impulse response filtering), elimination of unrelated variation (orthogonal signal correction)
fonollosa calibration transfer and drift counteraction direct standardisation 2016.pdf' | 2016 | Chemical Sensor Arrays | DS | Peer Reviewed academic paper | MOx gases | Inter-instrument | Aim is to counteract sensor sensitivity drift due to ageing of the sensors, to reduce time between recalibration. 5 sensing units over 22 days, 4 different MOx at 10 different concentrations. The mathematical model was based on SVM. DS recovered performance compared with unstandardised
Fonseca Diaz Unsupervised dynamic orthogonal projection. calibration transfer without standard 2022.pdf' | 2022 | NIRS | DOP | Peer Reviewed academic paper | Mange | Inter-instrument | Full text requested.
'Forina Transfer of Calibration NIRS 1995.pdf' | 1995 | NIRS | PDS | Peer Reviewed academic paper | Milled soy bean | Inter-instrument | 4 different laboratories each with a unique NIR instrument. Explains math theory extensively. Not well presemted, but conclusion was that PDS helped and only 2 PCs appear to be needed for the PDS step
'Guo Cal Transfer 3 soil nutrients VIS NIRS 2021.pdf' | 2021 | Vis-NIR | Variable selection, transformations | Peer Reviewed academic paper | Soil Organic Content | Output Standardisation | Not actual calibration transfer. Some useful exploration of preprocessing and variable selection procedures which can simplify transfer.
HUANG Chapter17-Calibrationtransfermethods 2022.pdf' | 2022 | General | Broad range | Book Chapter | General | Inter-instrument | Covers a lot of ground and goes into the theory of a number of methods in depth. Best used to look up theory behind about specific methods
'Li Standard free calibration transfer strategy discriminant model of apple origins NIRS 2022.pdf' | 2022 | VIS-NIR | PDS, SST, CTCCA | Peer Reviewed academic paper | Apple Soluble Solids | Inter-instrument | PDS performed best. S/B after PDS improved correlation but worsened the SEP
'Li calibration transfer portable VISNIR devices soluble solids apple 2022.pdf' | 2022 | NIRS | TCA, BDA, MEDA | Peer Reviewed academic paper | Apple Origin | Inter-instrument | Intended to rpovide solution where standards are not stable enough to run on multiple instruments. Tested three transfer learning algorithms , Transferred Component Analysis (TCA), Balanced Distribution Adaption (BDA), Manifold Embedded Distribution Alignment (MEDA). In 2021 described NIRS as still one model per one instrument. TCA and BDA had little impact, MEDA brought the accuracy of the second instruments close to that of 1st. Removing the wavelengths with the highest deviations ( a process called REA, relative error analysis) improved the performance of all three transfer methods, but MEDA remained the best.
'Mishra are standards needed for calibration transfer 2021.pdf' | 2021 | General | Wide range | Peer Reviewed academic paper, review | Non specific | Inter-instrument | Good summary of many techniques, useful reference to start exploring techniques of interest.
'Mou Robust Calibration Transfer 2016.pdf' | 2016 | IR | Cauchy (aka Lorentzian) | Peer Reviewed academic paper | General | Inter-instrument | This paper deals with the term used to deciede when error is minimised, rather than the algorithn ti transfer. It argues that the cauchy distribution more accurately potrays the distribution of variations between instruments than the more usual normal distribution (standard errors). Note that this may be technique specific as IR spectra are4 more likely to be lorentzian in shape than techniques with broad bands. Lorentzian occurs when indivdual units of variation positively influence each other, leading to a sharper central peak but broader tails (often labelled as outliers by SE methods).
'Nicolai Nondestructive_measurement_NIRS 2007.pdf' | 2007 | NIRS | review | Peer Reviewed academic paper | Fruit and Veg SSC | Inter-instrument, temperature | Gives overview of methods employed up to date of publication for using NIRS to measure SSC of various fruit and veg. Providers summary of standard pre-processing for SSC determination by NIRS and specifically mentions PDS and wavelets
'Nikzad-langerodi graph based calibration transfer 2020.pdf' | 2020 | NIRS | Manifold/graph | Peer Reviewed academic paper | cornmeal. publically available dataset | Inter-instrument | Along the lines of MEDA I think. Aim is to create a generalised transfer between instruments that is not so precisely linked to the application samples. It aims to do this by identifying the instrument specific deviations between a wide range of samples, and explicitly eliminate these. Most other methods instead explicitly look for the consistent features and actively model them. Used NBS glass standards
Panchuk calibration transfer between different analytical methods 2017.pdf' | 2017 | NIRS, XRF, UV-Vis | DS | Peer Reviewed academic paper | Sugar content Marzipan, metal ion concentrations | Inter-methodology | DS is able to allow transfer of calibrations between compltetely different methods.
'PerkinElmer application note MIRS 2024.pdf' | 2024 | IR | S/B, model updating | Technical White Paper | Biodiesel | Inter-instrument | Using ASTM standard method for calibrating is resource heavy. S/B and model updating are tested. Does not specify method used to calculate the transfer, but since it updates model it means no spectra are modified. Found that S/B was enough for this application and performed better than model updating
PM - 001 - 1 Model development for forages using NeoSpectra Instrument.docx' | 2024 | NIRS | Spiked Model | Technical Notes | forage | Inter-instrument, environmental, seasonal | Documentation from Si-ware on NeoSpectra system calibration. Mostly focused on validation, breif discussion of model maintenance. Advises data spiking to update the model and overtime remove original system data from the calibration set.
PM – 004 - 1 Quick Guide Spectral - Calibration Transfer.docx' | 2024 | NIRS | PDS | Technical Notes | General | Inter-instrument, environmental, seasonal | Recommends sample specific transfer function. Advises running samples same day, with identical settings. Samples should cover all expected variations. Need a validation set to test the transfer. Uses Eigenvector software ti calcuate function, recommends PDS
PM – 005 - 1 Model Transfer – Workflow.pptx' | 2024 | NIRS | Spiked Model | Technical Notes | General |  | Recommends at least 20% of spectra in calibration set are recorded on new instrument. Recommends transferring data from old system to new then rebuilding the calibration using a combination of the new instruiment and the transferred old data. This means an updated spiked model. Spectra will not need to be modified for the analysis.
Poerio Dual-Domain Calibration Transfer Using Orthogonbal Projection 2018.pdf' | 2018 | NIRS | wavelet then DOP | Peer Reviewed academic paper | several publically available datasets | Inter-instrument | Text is in DMR protected file. Recommends transfer in the frequency domain (i.e. wavlet or FT) saying it gives a more generalisable result. In essence dor DOP a matrix of spectral differences is created and orthogonally decomposed then these are projected on the data. This allows reconstruction of the differing elements and their subsequent elimination. Found that doing this in the frequency domain removed more irrelevant differences than in the spectral domain.
'Pu calibration transfer micro NIRS to hyperspectral imaging bananito_fruit 2018.pdf' | 2018 | NIRS, micro to imaging | PDS | Peer Reviewed academic paper | Fruit SSC | Inter-methodology | Preprocessing with SNV improved spectral differences from 31 to 9%. Original NIR model was R2 0.922, RMSEP 1.451%, NIR-HIS model with PDS was R2 0.925, RMSEP 1.592%. Found that spectral standardisation (SNV) was necessary before PDS
'Puleio Calibration of spectra non-stationary background unsupervised physics-informed deep learning 2023.pdf' | 2023 | general | N/A | Peer Reviewed academic paper | General theory | N/A | Not actual calibration transfer, but kept in as it does propose a novel AI based method to data pre-processing. The paper is very early stage PoC with simple challenges, but may be of some interest
'Rudnitskaya Calibration Update and Drift Correction E noses and tongues 2018.pdf' | 2018 | Chemical Sensor Arrays | Review | Peer Reviewed academic paper | Review | Inter-temporal | In cases where cause of drift is known and can be modelled from an independent variable (e.g. temperature, humidity) then this can be exploited to correct for those sources of drift. Drift correction relies of drift having common patterns, allowing methods such as PCA to separate out those patterns and be exploited to correct for them. OSC is a variant that eliminates patterns unrelated to the Y variable. ICA allows seperation of positive patterns, including drift patterns. Other approaches ar based on the frequency of the draift (low frequency, changes slowly over time) compared to analyte (changes with each sample, much more frequently), such as wavelet transforms. Time series analysis methods are also used such as ARMA, Kalman filter and chaotic time series. Downside is the need to monitor over long times to develop the appropriate models. Data standardisation methods discussed include S-B, PDS, Single Wave Standardisation (SWS). Also discuss mapping repsonses using ANNs and linear regression. Model expansion is a class of calibration transfer that involves adding new calibration standards that capture the new variation required and recalculating the model exposed to this expanded variation. Weighting is the simplest expansion method involving simply adding the new standards and recalulating the model weights. Tikhonov regularization is a variant of ridge regression, but where regularisation is aimed at improving calibration transfer. Joint -Y PLS allows for transfer of non-overlapping calibration sets with the requirement that the same number of variables are assessed in each original Y matrix and the X matrices share a covariance structure. Also discusses selection for the most efficicent set of samples for calibration transfer, e.g. using knowledge of relative concentrations, choosing high leverage samples, or Kennar-Stone algorithm (select uniform sdistribution over possibility space). The aim is to describe the extremes of variation along each important dimension of variation.
'Saeys multivariate calibration transfer Chemometrics Review_2019.docx' | 2019 | vibrational Spectroscopy, mostly NIRS | Review | Peer Reviewed academic paper | Fruit SSC | various | First step is to start with a calibration set that captures as much variation as possible (sources, cultivars, season). Putliers need to be handled appropriately [they are samples that are too underpowered to provide reliable statistical insight into that source of variation, making the model unstable]. Present useful summary of a range of multivariate modelling approaches for building a model. Discuss pre-processing for standardising signals before analysis. This includes transformations to a variable space with better mathematical properties (e.g. transmittance or reflectance converted to absobance), scatter correction (MSC, SNV, EMSC). Baseline correction discussed for Raman - mentions determining this on log(1/r) transform [this should be -Log] . Derivatives remove broad bandwidth curvature, but increase high frequency variability (e.g. shot noise). interference removal is discussed with OSC and OPLS (using just information within calibration set), External Parameter Orthogonalisation ( having a set of samples that specifically vary the interferent in order to explicity build a model to remove its contribution) or exploiting EMSC. Good in depth discussion of sample selection for calibration and the value of uniform sampling (uniform power across possibility space) as opposed to random (centres areround mean, low power for extremes, outliers harder to handle). Final independent validation on teh other hand should be representive of the expected distribution in use. Actual discussion of calibration transfer itself is very high level and vague. Useful overview of good practise in optimising within-instrument
'Sjoblom OSC calibration transfer NIRS 1998.pdf' | 1998 | NIRS | Standardisation, PDS, OSC | Peer Reviewed academic paper | Pharmaceutical in water | inter-instrument | Found that building a model based on both instruments needed no independent alignment methods. Main focus was on OSC but it did not appear to be as effective as some other approaches studied and provided no benefit when using the mdoel based on both instruments. The theory was that OSC would eliminate device specific variation irrelevant to target parameter.
Surkova CalTrans-LED-OMS-ACS-Sensors 2020.pdf' | 2020 | LED-optical multisensor | DS, S-B | Peer Reviewed academic paper | chemical in water | inter-instrument | LED-OMS are low cost low spec devices made of poorly standardised parts, making variability between devices high. Use of specific LEDs to illuminate at specific narrow bands of wavelength is the hardware version of variable selection. Looked at transfering calibrationfrom 2500 wavelength laboratory spectrometer to a 4 LED OMS. PDS not possible due to discontinuity between LEDs. Used DS to rebuild signals as if from lab instrument. Found that DS was effective, but no clear explanation why wanted to recreate full spectra, rather than just standardise the 4 wavelengths. 4 wavelengths of lab instrument gave same performance as full spectrum.
USRE34070 Regimand tarsnferring calibration data between calibrated instruments.pdf' | 1992 | N/A |  | Patent Filing |  |  | Method that uses cross correlation combined with daily backrgound readings to adjust offset for response in low quality field instruments.
US4864842 Regimand transferring calibration data between calibrated instruments.pdf | 1989 | N/A |  | Patent Filing |  |  | original uncorrected version (minor changes)
US4866644 Shenk Westerhaus Calibration Transfer.pdf | 1989 | N/A | S-W | Patent Filing | N/A | inter-instrument | patent filing for Shenk-Westerhaus method. Carries out a transfer for each variable in the signal. First the relationship between channels of the instruments is established, then the intensities are transformed by interpolating the channels found to be most closely related. [now superseeded by splines and interpolating to a standard channel pattern]
US5459677 Kowalski PDS.pdf | 1995 | N/A | PDS | Patent Filing | N/A | inter-instrument | Patent filing for PDS, using local pls models for each channel in signal. Since each channel in a spectra is correlated then information from surrounding channels can help refine the estimate of what the current channel should be.
US10082600B2 Shen calibration transfer for testing instrument.pdf | 2018 | FTIR | Extended Model? | Patent Filing | Fluid Analysis, competitive resevoir formation | inter-instrument | outlier rejection, splines to interpolate channels, baseline subtraction, calculate fractional transmittance, convert to absorbance, interference filters leading to clustering algorithim. Appears to be based around the specific application, rather than a novel algorithm.
US10429240B2 Hsiung Calibration transfer sparse.pdf | 2019 | N/A | Sample Selection | Patent Filing | N/A | inter-instrument | Sparse transfer set, appears to be sample selection procedure for creating a minimal sample set for efficicent transfer? Specifies the use of support vectors, I guess because this is the novel invention as sample selection was prior art.
US10969331B2 Sun Linear model correction for multivariate model maintenance.pdf | 2021 | N/A | Model Updating | Patent Filing | N/A | inter-instrument | Calculate model transfer coefficients to update model to handle new instrument. Uses two sets of coefficicents [presume to meet in the middle between the original and new systems?] Constraints used to guide it are minimising the difference between corrected instruments, requiring correlation between transformed coefficicents and the original, requriing slope between coefficicents and ensuring the final predictions is within the same dynamic range as the original.
US11378452B2 Hsiung Sparse calibration transfer.pdf | 2022 | N/A | Sample Selection | Patent Filing | N/A | inter-instrument | Appears to be using the earlier patent by the same author, but specifying that the transfer model is being created?
US20160327389A1 Uzunovic calibration transfer between two devices.pdf | 2016 | Compass | Model Updating | Patent Filing | Aerial drones | inter-instrument | Create calibration difference value file and use to udpate calibration
'Wang Improvement of multivariate calibration through instrument standardization 1992.pdf' | 1992 | NIR | PDS | Peer Reviewed academic paper | Fuel | inter-instrument | Found that PDS allowed a better prediction performance on instrument B than creating a new calibration directly on B. This was attributed to the original calibration being more thorough and on higher quality instrument.
'Wang Multivariate instrument standardization 1991.pdf' | 1991 | NIR | PDS, DS, inverse model standardisation, subset recalibration | Peer Reviewed academic paper | simulated data | inter-instrument | PDS was preferred method. Original PDS paper?
'Wang Standardization of second order instruments 1993.pdf' | 1993 | GC?MS LC/UV | dual band PDS | Peer Reviewed academic paper | hyphenated and 2D methods | inter-instrument | PDS for each dimension of the hyhenated method
'Workman Reference Materials I X axis Standards 2019.pdf' | 2019 | N/A | Instrument Standardisation | Trade publication | N/A | X-axis standardisation | Discussion of general principles of x-axis standards and technology speicific sections
'Workman Reference Materials II Photometric Standards 2019.pdf' | 2019 | N/A | Instrument Standardisation | Trade publication | N/A | Y-axis standardisation | Discussion of general principles of Y-axis standards and technology speicific sections
'Workman review-of-calibration-transfer-practices-and-instrument-differences-in-spectroscopy 2018.pdf' | 2018 | review | broad range | Peer Reviewed academic paper | review | inter-instrument | Review of calibration transfer methods and methods of testing instrument agreement
'Workman AI in analytical spectroscopy I Basic Concepts 2023.pdf' | 2023 | N/A | off topic | Trade publication | N/A | N/A | interesting discussion on AI for spectroscopists
'Workman AI in analytical spectroscopy II examples 2023.pdf' | 2023 | N/A | off topic | Trade publication | N/A | N/A | interesting discussion on AI for spectroscopists
'Workman Calibration Transfer Chemometrics I 2017.pdf' | 2017 | summary | PDS,DS,OSC,Procrustes, finite impulse response, maimum likelihood PCA | Trade publication | review | inter-instrument | Describes 4 main approaches: 1. instrument matching (make output of instruments uniform) 2. broad variation model (make model robust to differences) 3. model updating (recalulating original model with data from new system 4. select channels insensitive to differences. 5. adjusting the predictions by slope/bias adjustments. First step is instrument alignement and correction against standards
'Workman Calibration Transfer Chemometrics II 2018.pdf' | 2018 | summary | CCA, Spectral Regression, Compressed wavelet domain, wavelet packet transform standardisation, MSC, SNV, detrend, synthetic calibration spectra, Fourier transform transfer, temperature compensation, compensating for external phenomena | Trade publication | review | inter-instrument | continues part 1, giving overview of additional tools
'Workman Calibration Transfer I 2013.pdf' | 2013 | NIRS | Transforms | Trade publication | tutorial | inter-instrument | Discusses issues will calibration transfer when the calibration is not a correct one for the actual properties of the samples. Uses example of determining %wt values with NIRSwhose intensity varies by %Vol. Points out that calibration transfer of an inappropriate model will replicate that inappropriate behanviour.
'Workman Calibration Transfer II Instrumentation Aspects 2013.pdf' | 2013 | General | Standardisation | Trade publication | review | inter-instrument | Discusses ways to measure variations in performance between systems namely x-axis accuracy and repeatability, y-axis accuracy and repeatability, linearity of response, noise properties, signal averaging and inetnsity profile tests. Advise that carrying out these assessments can identify specific information for diagnoising causes of deviations. recommends calibrating instrument responses against agreed standards (e.g. ASTM, NIST)
'Workman Calibration Transfer III Maths Aspects 2013.pdf' | 2013 | General | Transfer Algorithms | Trade publication | review | inter-instrument | States that aligbning instrument response must be carried out first. Discusses S-B in details, mentions OSC for elimination of variation specific to instruments, model augmentation and mathematicl transformations
'Workman Calibration Transfer IV agreement between instruments 2013.pdf' | 2013 | General | Agreement Checking | Trade publication | review | inter-instrument | Focuses on methods to determine the level of agreement between instruments. Discusses uncertainty and how to deteermine it. Evaluates the Bland-Altman plot for evaluating bias relative to uncertainty
'Workman Calibration Transfer V Maths of wavelength standards 2014.pdf' | 2014 | NIRS | Standardisation | Trade publication | review | inter-instrument | Focusing on standardising the x-axis and issue with describing deviations in the x-axis calibration
'Workman Calibration Transfer VI Maths of Photometric standards 2014.pdf' | 2014 | General | Standardisation | Trade publication | review | inter-instrument | Discusses various causes of Y-axis response changes over time and between instruments. Discussess ways to align the Y-axis and to measure accuracy and precision and to compare uncertainyt caross instruments
'Xu calibration transfer optimized single kernel NIRS 2019.pdf' | 2019 | NIRS | Spectral Space Transformation | Peer Reviewed academic paper | Rice kernels | Sample preperation methods | Wanted to be able to measure dehusked single rice kernel to determine protein content. But wanted to make the analysis robust to sample preperation method (husk, dehusked and flour) so it would be more flexible. Standardised spectra using SNV. Used SST which appears quite similar in concept to DS. Was able to transfer the husked rice intoboth flour and husk calibration models, but direct application of each without the SST was poorer.
'Yan Calibration Transfer and drift compensation e noses via coupled task learning 2015.pdf' | 2015 | Electronic nose | coupled task learning | Peer Reviewed academic paper | Gases in breath | Temporal Drift, inter-instrument | Coupled Task Learning is a learning algorithm that shares coefficicents between related tasks, in this case the calibration against analyte and the drift in the sensors over time. A subset of samples is shared between models for the transfer. For the 2 sensor dataset 3-group classification was possible after CTL was applied. Optimal transferability achieved with only 10 transfer samples. On average it outperformed alternative methods for countering temporal drift
'Yang interpretable deep learning for calibration transfer NIRS 2022.pdf' | 2022 | NIRS | interpretable deep learning, PDS, CCA, S-B | Peer Reviewed academic paper | soybean and wheat | inter-instrument | Call the method DeepTranSpectra (DTS). Speifically design to transfer convolutional NN among multiple NIRS. Compared to PDS, CCA and S-B.5 different instrument models across 3 manufacturers (2 Foss, 2 Bruker, 1 Thermo). Interpolated to fixe wavenumber range and spacing by splines. The NN architecture is a series of connected feed forward netwrodks with convolutional layers reducing parameters between them. The apparoch first trains the parent set calibration and uses these to set the convolutional layers. Fine tuning is carried out on the child device. The DTS aims to use only samples on teh child for determining transfer - no parallel measurements. The transfer was signfiicanty improved compared to estalbised methods. inclusion of gradient weighted class activation mapping allowed visualisatio of which features were being used inthe transfer.
'Yap Calibration Transfer NIRS adaptive ANN Masters thesis 2021.pdf' | 2021 | NIRS | adaptive ANN (AANN) | Masters thesis | corn oil | Inter-instrument | Overall aim is to develop transfer approach when parent instrument no longer available. The models to be transferred were ANNs based on dat afrom parent. The ANN was then re-weighted (adaptive ANN, or AANN) using data collected on the child instrument only. This performed better than PDS-PLS.
'Yoon Calibration Transfer Algorithm for NIRS 2001.pdf' | 2001 | NIRS | OSC, 1st derivative, PDS | Peer Reviewed academic paper | benzene | temporal drift | Preprocessing by 1st derivative to remove baseline variation, OSC to remove variatio unrelaed to Y, Kennard Stone to select transfer samples then PDS to transfer. Found that the combination of 1st derivative and OSC were best at creating agreement before PDS was employed.Table 2 values quite a bot higher than Table 1 (parent calibration) so gain is not very clear
'Zhang New Algorithms for Calibration Transfer in Near Infrared Spectroscopy 2017.pdf' | 2017 | [5 empty] | Would need translation
'Zhao PLS subspace calibration transfer NIRS 2019.pdf' | 2019 | NIRS | PLS subspace | Peer Reviewed academic paper | corn, wheat, pharmaceutical | Inter-instrument | PLSCT involves projecting the parent PLS model onto the child data then OLS tarnsforms this into a representation of what it would be if measured on parent instrument. Simialr to DS, but using PLS latent variables instead of PCA ones. [Sounds like doing a PLS2 between the instruments?] [Strangely they use fewer LVs to transfer than their model requires. This means the child data has a lot of information wiped that the original calibration apparently needed. I am concerned about this]

## Abbreviations
NIRS | Near Infrared Spectroscopy
FT | Fourier Transform
PDS | Piecewise Direct Standardization
DS | Direct Standardisation
PAT | Process Analytical Technologies
EMSC | Extended Multiplicative Scattering Correction
MIRS | Mid-Infrared Spectroscopy
S-W | Shenk-Westerhaus
S/B | Slope Bias correction
ANN | artificial neural networks
SVM | Support Vector Machine
SST | spectral space transformation
CTCCA | calibration transfer canonical correlation analysis
TCA | Transferred Component Analysis
BDA | Balanced Distribution Adaption
MEDA | Manifold Embedded Distribution Alignment
TL | Transfer Learning
DOP | Dynamic orthogonal projection
REA | Relative Error Analysis
XRF | X-Ray Fluourescence
UV-Vis | UltraViolet-Visible Spectroscopy