## Slide 1
Transfer models between instruments
Applications Team
13-10-2022
Notes: v

## Slide 2
Objective
The main objective is to set the requirement for spectra transfer between instrument from different suppliers.
To have in mind:
Models cannot be transferred directly from one supplier to another
Only the spectra used for the calibration in the Inst. 1, can be transferred to the new instrument (Inst. 2)
There is a need to have native samples (20% of the Calibration matrix) to increase the performance of the models in final instrument

## Slide 3
Key points and phases
Key points:
Model to transfer the spectra from Inst. 1 to Inst. 2. This model would be only using spectral data from the instruments. (Not a constituent/s needed)
Once the transfer model is done, we can transfer the calibration spectra from Inst. 1 to Inst. 2.
Transferred spectra and native spectra from Inst. 2 (Neospectra) would be used to calibrate the constituents using PLSR.
The process is divided in 2 phases
Transfer model building
Predictive model building

## Slide 4
Transfer model building
Settings
20 samples of specific material are needed to build the model transfer between instruments. Samples should cover the calibration range of each parameter (no reference value is needed – Use the calibration on Inst. 1 to estimate the value)
Samples should be measured at the same time in both instruments:
Measurements on the Inst. 1 (old system) should be done the same way that the samples were collected on the calibration set of samples.
Measurements on the Inst. 2 (Si-Ware) should be done using the accessories, if any, for the final application.
20 samples
Inst. 1
Inst. 2
Instrument standardization model Building
Transfer model
(Inst.1 to Inst.2)

## Slide 5
Predictive model building
Settings
Once the transfer model is done, it can be applied to the calibration set from Inst. 1 (old)
Needs:
Calibration matrix used on the Inst. 1
20% of the native samples (Inst. 2 – Si-Ware) are needed to build a correct predictive model
Original Cal Inst. 1
Si-Ware samples 20% of the final Cal. matrix
Transfer model
(Inst.1 to Inst.2)
Transferred spectra Inst. 1 model Building
Parameter of interest
+
PLSR model
Remove spectral duplicate samples